## FP-ENG-05-01 rev.02
ทะเบียนควบคุมและประวัติการสอบเทียบเครื่องมือวัด | [13 empty] | เอกสารเลขที่ : FP-ENG-05-01 Rev.02
เริ่มใช้ :
เลขที่ | ชื่อเครื่องมือวัด | รุ่นเครื่องมือวัด | หมายเลขเครื่อง | จุดประสงค์ที่ใช้วัด | ช่วงในการวัดของเครื่องมือวัด | ค่าผิดพลาดของเครื่องมือวัด | ช่วงในการใช้งานของผลิตภัณฑ์ | เกณฑ์การยอมรับของผลิตภัณฑ์ | ผู้รับผิดชอบ | ความถี่ในการสอบเทียบ | สอบเทียบครั้งล่าสุด | สอบเทียบครั้งต่อไป | สถานที่ | สถานะ | หมายเหตุ
1 | kVp Meter | Piranha 557 | CB2-14060199 | kVp | 35-160kVp | ± 1.5 % | 40-125 kVp | ± 5 % | Service admin | 1 ปี | 2026-04-03 | 2027-04-03 | Head office | ปกติ
|  |  | CB2-14060199 | Exposure Time | 0.1ms-2000s | ± 1.0 % | 1msec-6sec | ± 10 % | Service admin | 1 ปี | 2026-04-03 | 2027-04-03
2 | Digital Pressure Gauge | N/A | AE969 | Pressure | 0-1500 psi | 1.8 psi | 0-225 psi | ± 50 psi | Service admin | 1 ปี | 2024-01-17 | 2025-01-17 | Head office | ยกเลิกการใช้งาน 1/25
3 | Pressure Calibrator | Fluke-717 30G | 3743662 | Pressure | 0-30 psi | ± 0.12 psi | 0-15 psi | ± 1 psi | Service admin | 1 ปี | 2026-06-12 | 2027-06-12 | Head office | ปกติ
4 | Pressure Calibrator | Fluke-717 30G | 5401455 | Pressure | 0-30 psi | ± 0.12 psi | 0-15 psi | ± 1 psi | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office | ปกติ
5 | Pressure Calibrator | Fluke-717 30G | 5401456 | Pressure | 0-30 psi | ± 0.12 psi | 0-15 psi | ± 1 psi | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office | ปกติ
6 | Electrical Safety Analyzer | ESA620 | 2629220 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2026-06-01 | 2027-06-01 | QA Team | ปกติ
|  | ESA620 | 2629220 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2026-06-01 | 2027-06-01 | QA Team
|  | ESA620 | 2629220 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2026-06-01 | 2027-06-01 | QA Team
|  | ESA620 | 2629220 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2026-06-01 | 2027-06-01 | QA Team
7 | Electrical Safety Analyzer | ESA612 | 4554080 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office | ปกติ
|  | ESA612 | 4554080 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | ESA612 | 4554080 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | ESA612 | 4554080 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
8 | Electrical Safety Analyzer | ESA612 | 4896023 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2025-07-01 | 2026-07-01 | Head office | ส่งแคล 19 Jun 2026
|  | ESA612 | 4896023 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2025-07-01 | 2026-07-01 | Head office
|  | ESA612 | 4896023 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2025-07-01 | 2026-07-01 | Head office
|  | ESA612 | 4896023 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2025-07-01 | 2026-07-01 | Head office
9 | Electrical Safety Analyzer | ESA612 | 4985040 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2025-07-14 | 2026-07-14 | Chiang Mai Branch | ส่งแคล 7 July 2026
|  | ESA612 | 4985040 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2025-07-14 | 2026-07-14 | Chiang Mai Branch
|  | ESA612 | 4985040 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2025-07-14 | 2026-07-14 | Chiang Mai Branch
|  | ESA612 | 4985040 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2025-07-14 | 2026-07-14 | Chiang Mai Branch
10 | Electrical Safety Analyzer | ESA612 | 5040061 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2025-08-11 | 2026-08-11 | KhonKaen Branch | ปกติ
|  | ESA612 | 5040061 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2025-08-11 | 2026-08-11 | KhonKaen Branch
|  | ESA612 | 5040061 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2025-08-11 | 2026-08-11 | KhonKaen Branch
|  | ESA612 | 5040061 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2025-08-11 | 2026-08-11 | KhonKaen Branch
11 | Electrical Safety Analyzer | ESA612 | 5040062 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2025-06-20 | 2026-06-20 | Hat Yai Branch | ปกติ
|  | ESA612 | 5040062 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2025-06-20 | 2026-06-20 | Hat Yai Branch
|  | ESA612 | 5040062 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2025-06-20 | 2026-06-20 | Hat Yai Branch
|  | ESA612 | 5040062 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2025-06-20 | 2026-06-20 | Hat Yai Branch
12 | Electrical Safety Analyzer | ESA612 | 5373501 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office | ปกติ
|  | ESA612 | 5373501 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
|  | ESA612 | 5373501 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
|  | ESA612 | 5373501 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
13 | Electrical Safety Analyzer | ESA612 | 5373504 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office | ปกติ | K.Anupong
|  | ESA612 | 5373504 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
|  | ESA612 | 5373504 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
|  | ESA612 | 5373504 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2026-05-06 | 2027-05-06 | Head office
14 | Fetal Simulator | PS320 | 5418028 | US (BPM) | 30-240 BPM | ±1 BPM | 50-210 BPM | ±1 BPM | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office | ปกติ
|  | PS320 | 5418028 | UA(mmHg) | 0-100 | ±1 mmHg | 10-90 mmHg | ±1 mmHg | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
15 | Defibrillator/Transcutanreous Analyzer | Impulse7000DP | 2986072 | Energy | 0.1-600 Joules | ±(1% of Reading +0.1J) | 2 -360 Joule | ± 10 % of Setting | Service admin | 1 ปี | 2025-10-27 | 2026-10-27 | Head office | ปกติ
16 | Defibrillator/Transcutanreous Analyzer | Impulse7000DP | 5383046 | Energy | 0.1-600 Joules | ±(1% of Reading +0.1J) | 2 -360 Joule | ± 10 % of Setting | Service admin | 1 ปี | 2026-03-19 | 2027-03-19 | Head office | ปกติ
17 | Gas Flow Analyzer | VT650 | 5360407 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office | ปกติ
|  | VT650 | 5360407 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
|  | VT650 | 5360407 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
18 | Gas Flow Analyzer | VT650 | 5361070 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2026-04-10 | 2027-04-10 | Hatyai Branch | ปกติ
|  | VT650 | 5361070 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2026-04-10 | 2027-04-10 | Hatyai Branch
|  | VT650 | 5361070 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2026-04-10 | 2027-04-10 | Hatyai Branch
19 | True RMS Multimeter | Fluke-116 | 58911057WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office(Ratthammanoon) | ปกติ
|  | Fluke-116 | 58911057WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 58911057WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 58911057WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 58911057WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
20 | Infusion Analyzer | IDA-1S | 5367062 | Flow Rate | 0.5 to 1000ml/h | 1% of Reading | 1-999 ml/h | ± 5 % | Service admin | 1 ปี | 2025-12-01 | 2026-12-01 | Head office | ปกติ
|  | IDA-1S | 5367062 | Volume | 0.06 to 999 ml | 1% of Reading | 1-999 ml | ± 10 % | Service admin | 1 ปี | 2025-12-01 | 2026-12-01 | Head office
21 | Infusion Analyzer | IDA-1S | 5471034 | Flow Rate | 0.5 to 1000ml/h | 1% of Reading | 1-999 ml/h | ± 5 % | Service admin | 1 ปี | 2025-12-01 | 2026-12-01 | Head office | ปกติ
|  | IDA-1S | 5471034 | Volume | 0.06 to 999 ml | 1% of Reading | 1-999 ml | ± 10 % | Service admin | 1 ปี | 2025-12-01 | 2026-12-01 | Head office
22 | BilBlanket Light Meter | Ohmeda Medical | 41001991 | Lighting | 5-60 µW/cm²/nm | 0.15 µW/cm²/nm | 20-50 µW/cm²/nm | ± 25 % | Service admin | 1 ปี | 2025-11-24 | 2026-11-24 | Head office | ปกติ
23 | Digital Thermometer withThermocople Type K | Fluke 51 II with Probe | 36720143WS,None | Temperature | 20-40 °C | ±0.2 °C | 30-40 °C | ±0.3 °C | Service admin | 1 ปี | 2026-03-18 | 2027-03-18 | Head office | ปกติ
24 | Incubator/Radiant Warmer Analyzer | INCU II | 54080004 | Temperature | 0-50 °C | ±0.05 °C | 30-42 °C | ±0.3 °C | Service admin | 1 ปี | 2025-08-26 | 2026-08-26 | Head office | ปกติ
25 | Vital Signs Simulator | Prosim4 | 4205029 | Heart rate | 30-320 BPM | ±1% of Setting | 20-300 BPM | ± 5 BPM | Service admin | 1 ปี | 2024-11-26 | 2025-11-26 | Head office | ส่งแคล 19 Jun 2026
|  | Prosim4 | 4205029 | Non invasive blood pressure | 10-400mmHg | ±(1% of reading +1mmHg) | 30-290 mmHg | ± 5 mmHg | Service admin | 1 ปี | 2024-11-26 | 2025-11-26 | Head office
|  | Prosim4 | 4205029 | Invasive blood pressure | 0-250 mmHg | ±(1% of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2024-11-26 | 2025-11-26 | Head office
|  | Spot Light | 4205029 | Percent Oxygen | 70-100 % | ±(1+Accuracy of UUT) | 90-99 % | ± 3 % | Service admin | 1 ปี | 2024-11-26 | 2025-11-26 | Head office | ส่งแคล 19 Jun 2026
26 | Vital Signs Simulator | Prosim4 | 4288029 | Heart rate | 30-320 BPM | ±1% of Setting | 20-300 BPM | ± 5 BPM | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office | ปกติ
|  | Prosim4 | 4288029 | Non invasive blood pressure | 10-400mmHg | ±(1% of reading +1mmHg) | 30-290 mmHg | ± 5 mmHg | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | Prosim4 | 4288029 | Invasive blood pressure | 0-250 mmHg | ±(1% of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | Spot Light | 5670083 | Percent Oxygen | 70-100 % | ±(1+Accuracy of UUT) | 90-99 % | ± 3 % | Service admin | 1 ปี | 2026-06-01 | 2027-06-01 | Head office
27 | Vital Signs Simulator | Prosim8 | 3556046 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-08-06 | 2026-08-06 | QA Team | ปกติ |  | TE28000049
|  | Prosim8 | 3556046 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-08-06 | 2026-08-06 | QA Team
|  | Prosim8 | 3556046 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-08-06 | 2026-08-06 | QA Team
|  | Prosim8 | 3556046 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-08-06 | 2026-08-06 | QA Team
|  | Prosim8 | 3556046 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-08-06 | 2026-08-06 | QA Team
28 | Vital Signs Simulator | Prosim8 | 4689688 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Head office | ปกติ
|  | Prosim8 | 4689688 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Head office
|  | Prosim8 | 4689688 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Head office
|  | Prosim8 | 4689688 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Head office
|  | Prosim8 | 4689688 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Head office
29 | Vital Signs Simulator | Prosim8 | 5039036 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | KhonKaen Branch | ส่งแคล 7 July 2026
|  | Prosim8 | 5039036 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | KhonKaen Branch
|  | Prosim8 | 5039036 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | KhonKaen Branch
|  | Prosim8 | 5039036 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | KhonKaen Branch
|  | Prosim8 | 5039036 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | KhonKaen Branch
30 | Vital Signs Simulator | Prosim8 | 5019606 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Hat Yai Branch | ปกติ
|  | Prosim8 | 5019606 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Hat Yai Branch
|  | Prosim8 | 5019606 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Hat Yai Branch
|  | Prosim8 | 5019606 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Hat Yai Branch
|  | Prosim8 | 5019606 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Hat Yai Branch
31 | Vital Signs Simulator | Prosim8 | 5019604 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Chiang Mai Branch | ส่งแคล 7 July 2026
|  | Prosim8 | 5019604 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Chiang Mai Branch
|  | Prosim8 | 5019604 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Chiang Mai Branch
|  | Prosim8 | 5019604 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Chiang Mai Branch
|  | Prosim8 | 5019604 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-07-16 | 2026-07-16 | Chiang Mai Branch
32 | Vital Signs Simulator | Prosim8 | 5040076 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Central | ปกติ |  | TE28000125
|  | Prosim8 | 5040076 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Central
|  | Prosim8 | 5040076 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Central
|  | Prosim8 | 5040076 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Central
|  | Prosim8 | 5040076 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2025-10-24 | 2026-10-24 | Central
33 | Vital Signs Simulator | ProSim8 | 5412040 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2026-06-04 | 2027-06-04 | Head office | ปกติ
|  | Prosim8 | 5412040 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2026-06-04 | 2027-06-04 | Head office
|  | Prosim8 | 5412040 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2026-06-04 | 2027-06-04 | Head office
|  | Prosim8 | 5412040 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2026-06-04 | 2027-06-04 | Head office
|  | Prosim8 | 5412040 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2026-06-04 | 2027-06-04 | Head office
34 | Vital Signs Simulator | ProSim8 | 5411059 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2026-06-02 | 2027-06-02 | Head office | ปกติ |  | TE28000169
|  | ProSim8 | 5411059 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2026-06-02 | 2027-06-02 | Head office
|  | ProSim8 | 5411059 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2026-06-02 | 2027-06-02 | Head office
|  | ProSim8 | 5411059 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2026-06-02 | 2027-06-02 | Head office
|  | ProSim8 | 5411059 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2026-06-02 | 2027-06-02 | Head office
35 | True RMS Multimeter | Fluke-115 | 37432863WS | AC Current | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office ( Witsawa ) | ปกติ
|  | Fluke-115 | 37432863WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 37432863WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 37432863WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 37432863WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
36 | True RMS Multimeter | Fluke-115 | 37432964WS | AC Current | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office (Thongchai ) | ปกติ
|  | Fluke-115 | 37432964WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432964WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432964WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432964WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
37 | True RMS Multimeter | Fluke-115 | 30702177WS | AC Current | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-04-23 | 2027-04-23 | Head office(K.Akkapol) | ปกติ
|  | Fluke-115 | 30702177WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-04-23 | 2027-04-23 | Head office
|  | Fluke-115 | 30702177WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-04-23 | 2027-04-23 | Head office
|  | Fluke-115 | 30702177WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-04-23 | 2027-04-23 | Head office
|  | Fluke-115 | 30702177WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-04-23 | 2027-04-23 | Head office
38 | True RMS Multimeter | Fluke-115 | 37432903WS | AC Current | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office(Khongpan) | ปกติ
|  | Fluke-115 | 37432903WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 37432903WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 37432903WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 37432903WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
39 | True RMS Multimeter | Fluke-115 | 37432905WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office(Head Office) | ปกติ
|  | Fluke-115 | 37432905WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432905WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432905WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432905WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
40 | True RMS Multimeter | Fluke-115 | 30702180WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office(K.Itsara) | ปกติ
|  | Fluke-115 | 30702180WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 30702180WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 30702180WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 30702180WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
41 | True RMS Multimeter | Fluke-115 | 39672295WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office(K.Chankrit) | ปกติ
|  | Fluke-115 | 39672295WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 39672295WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 39672295WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 39672295WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
42 | True RMS Multimeter | Fluke-115 | 40293201WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office (Chaiwoot) | ปกติ
|  | Fluke-115 | 40293201WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293201WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293201WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293201WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
43 | True RMS Multimeter | Fluke-115 | 40293203WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office (Aeknachai) | ปกติ
|  | Fluke-115 | 40293203WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293203WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293203WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293203WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
44 | True RMS Multimeter | Fluke-115 | 37432904WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office (Nonthawat ) | ปกติ
|  | Fluke-115 | 37432904WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432904WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432904WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432904WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
45 | True RMS Multimeter | Fluke-115 | 39672294WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office (Pongsak) | ปกติ
|  | Fluke-115 | 39672294WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 39672294WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 39672294WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
|  | Fluke-115 | 39672294WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-12-15 | 2026-12-15 | Head office
46 | True RMS Multimeter | Fluke-115 | 27800389WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office(Pattarapon) | ปกติ
|  | Fluke-115 | 27800389WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 27800389WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 27800389WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 27800389WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
47 | True RMS Multimeter | Fluke-115 | 34793538WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office(K.Pumarin) | ปกติ
|  | Fluke-115 | 34793538WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 34793538WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 34793538WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 34793538WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
48 | True RMS Multimeter | Fluke-115 | 37432865WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office ( Nawapol ) | ปกติ
|  | Fluke-115 | 37432865WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432865WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432865WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
|  | Fluke-115 | 37432865WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-12-04 | 2026-12-04 | Head office
49 | True RMS Multimeter | Fluke-115 | 27800390WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office ( Nattapol ) | ปกติ
|  | Fluke-115 | 27800390WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 27800390WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 27800390WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 27800390WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
50 | True RMS Multimeter | Fluke-115 | 34793511WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office หวาย | ปกติ
|  | Fluke-115 | 34793511WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 34793511WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 34793511WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 34793511WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
51 | True RMS Multimeter | Fluke-115 | 37432901WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office (Yuttana) | ปกติ
|  | Fluke-115 | 37432901WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432901WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432901WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
|  | Fluke-115 | 37432901WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-28 | 2026-11-28 | Head office
52 | True RMS Multimeter | Fluke-115 | 47580557SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office (Chayawich) | ปกติ
|  | Fluke-115 | 47580557SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580557SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580557SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580557SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
53 | True RMS Multimeter | Fluke-115 | 47580560SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office (Suradet) | ปกติ
|  | Fluke-115 | 47580560SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580560SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580560SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580560SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
54 | True RMS Multimeter | Fluke-115 | 47580558SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office (Naruebet) | ปกติ
|  | Fluke-115 | 47580558SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580558SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580558SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
|  | Fluke-115 | 47580558SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-11-10 | 2026-11-10 | Head office
55 | True RMS Multimeter | Fluke-115 | 40293202WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office(K.Anupong) | ปกติ
|  | Fluke-115 | 40293202WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293202WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293202WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 40293202WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
56 | True RMS Multimeter | Fluke-115 | 37432864WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office(Niwat) | ปกติ
|  | Fluke-115 | 37432864WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432864WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432864WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
|  | Fluke-115 | 37432864WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-02-24 | 2027-02-24 | Head office
57 | True RMS Multimeter | Fluke-115 | 47170269SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office(K.Kergrit) | ปกติ
|  | Fluke-115 | 47170269SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
|  | Fluke-115 | 47170269SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
|  | Fluke-115 | 47170269SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
|  | Fluke-115 | 47170269SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office
58 | True RMS Multimeter | Fluke-115 | 47170268SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office(K.Chatmongkol) | ปกติ
|  | Fluke-115 | 47170268SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 47170268SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 47170268SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 47170268SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
59 | True RMS Multimeter | Fluke-115 | 38623123WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office(K.Jidnipid) | ปกติ
|  | Fluke-115 | 38623123WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 38623123WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 38623123WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 38623123WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
60 | True RMS Multimeter | Fluke-115 | 47170267SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office(K.Siripan) | ปกติ
|  | Fluke-115 | 47170267SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 47170267SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 47170267SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-115 | 47170267SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
61 | True RMS Multimeter | Fluke-115 | 47170266SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office (K.Anupong(Kat)) | ปกติ
|  | Fluke-115 | 47170266SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 47170266SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 47170266SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 47170266SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
62 | True RMS Multimeter | Fluke-115 | 58212165WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office (Yodsaphat) | ปกติ
|  | Fluke-115 | 58212165WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 58212165WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 58212165WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 58212165WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
63 | Gas Flow Analyzer | VT650 | 5610023 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Chiang Mai Branch | ปกติ
|  | VT650 | 5610023 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Chiang Mai Branch
|  | VT650 | 5610023 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Chiang Mai Branch
64 | Gas Flow Analyzer | VT650 | 5614085 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2025-09-29 | 2026-09-29 | KhonKaen Branch | ปกติ
|  | VT650 | 5614085 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2025-09-29 | 2026-09-29 | KhonKaen Branch
|  | VT650 | 5614085 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2025-09-29 | 2026-09-29 | KhonKaen Branch
65 | Vital Signs Simulator | Prosim8 | 5684042 | Heart rate | 10-300 BPM | ±1% of Setting | 20-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2026-02-16 | 2027-02-16 | QA Team | ปกติ
|  | Prosim8 | 5684042 | Non invasive blood pressure | 0-400 mmHg | ±(0.5%of reading +0.5mmHg) | 30-290 mmHg | ± 3 mmHg | Service admin | 1 ปี | 2026-02-16 | 2027-02-16 | QA Team
|  | Prosim8 | 5684042 | Invasive blood pressure | -10-300mmHg | ±(1%of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2026-02-16 | 2027-02-16 | QA Team
|  | Prosim8 | 5684042 | Percent Oxygen | 30-100 % | ± (1 + Accuracy of UUT) | 70-100 % | ± 2 % | Service admin | 1 ปี | 2026-02-16 | 2027-02-16 | QA Team
|  | Prosim8 | 5684042 | Temperature | 30-42 °C | ±0.4 °C | 31-40 °C | ±0.5 °C | Service admin | 1 ปี | 2026-02-16 | 2027-02-16 | QA Team
66 | ECG Simulator | MiniSim | 30355 | Heart rate | 30-300 BPM | ± 1 BPM | 30-300 BPM | ± 1 BPM | Service admin | 1 ปี | 2025-07-22 | 2026-07-22 | Head office | ส่งแคล 7 July 2026
67 | Decade Resistor | RS-200 | D2-042523139 | Resistance | 0-9 MΩ | ± (1%+0.025Ω) | 5Ω - 355Ω | ± 5 Ω | Service admin | 1 ปี | 2025-09-04 | 2026-09-04 | Head office | ปกติ
68 | Vital Signs Simulator | Prosim4 | 5408053 | Heart rate | 30-320 BPM | ±1% of Setting | 20-300 BPM | ± 5 BPM | Service admin | 1 ปี | 2025-05-27 | 2026-05-27 | Head office(15/3/23) | ส่งแคล 19 Jun 2026
|  | Prosim4 | 5408053 | Non invasive blood pressure | 10-400mmHg | ±(1% of reading +1mmHg) | 30-290 mmHg | ± 5 mmHg | Service admin | 1 ปี | 2025-05-27 | 2026-05-27 | Head office
|  | Prosim4 | 5408053 | Invasive blood pressure | 0-250 mmHg | ±(1% of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-05-27 | 2026-05-27 | Head office
|  | Spot Light | 6062036 | Percent Oxygen | 70-100 % | ±(1+Accuracy of UUT) | 90-99 % | ± 3 % | Service admin | 1 ปี | 2025-05-27 | 2026-05-27 | Head office
69 | Gas Flow Analyzer | VT650 | 5675106 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2026-06-12 | 2027-06-12 | QA Team | ปกติ
|  | VT650 | 5675106 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2026-06-12 | 2027-06-12 | QA Team
|  | VT650 | 5675106 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2026-06-12 | 2027-06-12 | QA Team
70 | Gas Flow Analyzer | VT650 | 5670104 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2026-05-15 | 2027-05-15 | Head office | ปกติ
|  | VT650 | 5670104 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2026-05-15 | 2027-05-15 | Head office
|  | VT650 | 5670104 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2026-05-15 | 2027-05-15 | Head office
71 | Vital Signs Simulator | Prosim4 | 6084028 | Heart rate | 30-320 BPM | ±1% of Setting | 20-300 BPM | ± 5 BPM | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Central(25/4/23) | ปกติ
|  | Prosim4 | 6084028 | Non invasive blood pressure | 10-400mmHg | ±(1% of reading +1mmHg) | 30-290 mmHg | ± 5 mmHg | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Prosim4 | 6084028 | Invasive blood pressure | 0-250 mmHg | ±(1% of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Spot Light | 6097001 | Percent Oxygen | 70-100 % | ±(1+Accuracy of UUT) | 90-99 % | ± 3 % | Service admin | 1 ปี | 2025-09-02 | 2026-09-02 | Head office
72 | Steel Weight | Class M1 | M1801082N(ID4)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | Head office | ยกเลิกการใช้งาน 2/26
73 | Steel Weight | Class M1 | M2103140S(ID1)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
74 | Steel Weight | Class M1 | M2103140S(ID2)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
75 | Steel Weight | Class M1 | M2103140S(ID3)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | Head office | ยกเลิกการใช้งาน 2/26
76 | Steel Weight | Class M1 | M2103140S(ID4)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
77 | Steel Weight | Class M1 | M2103140S(ID5)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
78 | Steel Weight | Class M1 | M2103140S(ID6)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
79 | Steel Weight | Class M1 | M2103140S(ID7)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
80 | Steel Weight | Class M1 | M2103140S(ID8)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2023-03-30 | 2024-03-30 | Head office | ยกเลิกการใช้งาน3/24
81 | Steel Weight | Class M1 | M1801082N(ID1)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | Head office | ยกเลิกการใช้งาน 2/26
82 | Steel Weight | Class M1 | M1801082N(ID3)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
83 | Steel Weight | Class M1 | M2103140S(ID1)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2023-03-30 | 2024-03-30 | Head office | ยกเลิกการใช้งาน3/24
84 | Steel Weight | Class M1 | M2104013N(ID1)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Head office | ปกติ
85 | Steel Weight | Class M1 | M2104013N(ID2)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2023-09-06 | 2024-09-06 | Head office | ยกเลิกการใช้งาน9/24
86 | Steel Weight | Class M1 | M2104013N(ID3)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
87 | Steel Weight | Class M1 | M2104013N(ID4)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2023-09-06 | 2024-09-06 | Head office | ยกเลิกการใช้งาน9/24
88 | Steel Weight | Class M1 | M2104013N(ID5)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ยกเลิกการใช้งาน9/24
89 | Steel Weight | Class M1 | M2104013N(ID6)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
90 | Steel Weight | Class M1 | M2104013N(ID7)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
91 | Steel Weight | Class M1 | M2112061N(ID1)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2025-12-16 | 2026-12-16 | Chiang Mai Branch | ปกติ
92 | Steel Weight | Class M1 | M2112061N(ID2)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2025-12-16 | 2026-12-16 | Chiang Mai Branch | ปกติ
93 | Steel Weight | Class M1 | M2112061N(ID3)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2023-01-12 | 2024-01-12 | Head office | ยกเลิกการใช้งาน1/24
94 | Steel Weight | Class M1 | M2112061N(ID4)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2023-01-12 | 2024-01-12 | Head office | ยกเลิกการใช้งาน1/24
95 | Steel Weight | Class M1 | M2112061N(ID5)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2024-01-09 | 2025-01-09 | Head office | ยกเลิกการใช้งาน1/25
96 | Steel Weight | Class M1 | M2112061N(ID6)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2024-01-09 | 2025-01-09 | Head office | ยกเลิกการใช้งาน1/25
97 | Steel Weight | Class M1 | M2112061N(ID7)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Hat Yai Branch | ปกติ
98 | Steel Weight | Class M1 | M2112061N(ID8)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Hat Yai Branch | ปกติ
99 | Steel Weight | Class M1 | M2003166N(ID3)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Workshop room | ปกติ
100 | Steel Weight | Class M1 | M2003166N(ID12)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
101 | Steel Weight | Class M1 | M2103140S(ID1)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
102 | Steel Weight | Class M1 | M2103140S(ID2)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | Head office | ยกเลิกการใช้งาน 2/26
103 | Steel Weight | Class M1 | M2103140S(ID3)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
104 | Steel Weight | Class M1 | M2103140S(ID4)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2023-03-30 | 2024-03-30 | Head office | ยกเลิกการใช้งาน3/24
105 | Steel Weight | Class M1 | M2103140S(ID5)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2024-05-21 | 2025-05-21 | Head office | ยกเลิกการใช้งาน5/25
106 | Steel Weight | Class M1 | M2103140S(ID6)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
107 | Steel Weight | Class M1 | M2103140S(ID7)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
108 | Steel Weight | Class M1 | M2103140S(ID8)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
109 | Steel Weight | Class M1 | M2103140S(ID9)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | Workshop room | ยกเลิกการใช้งาน3/24
110 | Steel Weight | Class M1 | M2103140S(ID10)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
111 | Steel Weight | Class M1 | M2103140S(ID11)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | QA Team | ยกเลิกการใช้งาน3/24
112 | Steel Weight | Class M1 | M2103140S(ID12)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
113 | Steel Weight | Class M1 | M2103140S(ID13)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
114 | Steel Weight | Class M1 | M2103140S(ID14)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
115 | Steel Weight | Class M1 | M2103140S(ID15)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
116 | Steel Weight | Class M1 | M2103140S(ID16)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
117 | Steel Weight | Class M1 | M2103140S(ID17)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 21/5/2025 | 21/5/2025 | QA Team | ยกเลิกการใช้งาน5/25
118 | Steel Weight | Class M1 | M2103140S(ID18)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
119 | Steel Weight | Class M1 | M2104013N(ID1)-10KG | Weight | 0-10kg | ± 160 mg | 0-10000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
120 | Steel Weight | Class M1 | M1711191N(ID2)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
121 | Steel Weight | Class M1 | M1801082N(ID2)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
122 | Steel Weight | Class M1 | M1906195N(ID1)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Workshop room | ปกติ
123 | Steel Weight | Class M1 | M1906195N(ID4)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | Workshop room | ยกเลิกการใช้งาน 2/26
124 | Steel Weight | Class M1 | M2003166N(ID1)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Head office | ปกติ
125 | Steel Weight | Class M1 | M2003166N(ID2)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
126 | Steel Weight | Class M1 | M2103140S(ID1)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | QA Team | ปกติ
127 | Steel Weight | Class M1 | M2103140S(ID2)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | QA Team | ยกเลิกการใช้งาน5/25
128 | Steel Weight | Class M1 | M2103140S(ID3)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2024-05-21 | 2025-05-21 | QA Team | ยกเลิกการใช้งาน5/25
129 | Steel Weight | Class M1 | M2103140S(ID4)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Workshop room | ปกติ
130 | Steel Weight | Class M1 | M2103140S(ID5)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ยกเลิกการใช้งาน5/25
131 | Steel Weight | Class M1 | M2103140S(ID6)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Workshop room | ปกติ
132 | Steel Weight | Class M1 | M2103140S(ID7)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
133 | Steel Weight | Class M1 | M2104013N(ID1)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | QA Team | ปกติ
134 | Steel Weight | Class M1 | M2104013N(ID2)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ยกเลิกการใช้งาน4/25
135 | Steel Weight | Class M1 | M2104013N(ID3)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | QA Team | ยกเลิกการใช้งาน 2/26
136 | Steel Weight | Class M1 | M2104013N(ID4)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | QA Team | ยกเลิกการใช้งาน3/24
137 | Steel Weight | Class M1 | M2104013N(ID5)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-02-27 | 2026-02-27 | QA Team | ยกเลิกการใช้งาน 2/26
138 | Steel Weight | Class M1 | M2104013N(ID6)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Head office | ปกติ
139 | Steel Weight | Class M1 | M2104013N(ID7)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | QA Team | ยกเลิกการใช้งาน3/24
140 | Steel Weight | Class M1 | M2104013N(ID8)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | QA Team | ยกเลิกการใช้งาน3/24
141 | Steel Weight | Class M1 | M2104013N(ID9)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
142 | Steel Weight | Class M1 | M2104013N(ID10)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
143 | Steel Weight | Class M1 | M2104013N(ID11)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2023-03-09 | 2024-03-09 | QA Team | ยกเลิกการใช้งาน3/24
144 | Steel Weight | Class M1 | M2104013N(ID12)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2025-08-07 | 2026-08-07 | Head office | ปกติ
145 | True RMS Multimeter | Fluke-116 | 47260492SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office(Sowasak) | ปกติ
|  | Fluke-116 | 47260492SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-116 | 47260492SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-116 | 47260492SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-116 | 47260492SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
146 | True RMS Multimeter | Fluke-115 | 27800417WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office(Woravit) | ปกติ
|  | Fluke-115 | 27800417WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-115 | 27800417WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-115 | 27800417WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-115 | 27800417WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
147 | True RMS Multimeter | Fluke-115 | 37432902WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office(Wongsakon) | ปกติ
|  | Fluke-115 | 37432902WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 37432902WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 37432902WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
|  | Fluke-115 | 37432902WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-03 | 2026-09-03 | Head office
148 | True RMS Multimeter | Fluke-116 | 48240758SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office(Takonkorn)
|  | Fluke-116 | 48240758SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-116 | 48240758SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-116 | 48240758SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-116 | 48240758SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
149 | True RMS Multimeter | Fluke-116 | 48920301SV | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office(Chalermpol) | ปกติ
|  | Fluke-116 | 48920301SV | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 48920301SV | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 48920301SV | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
|  | Fluke-116 | 48920301SV | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-16 | 2026-09-16 | Head office
150 | True RMS Multimeter | Fluke-179 | 91560315 | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office(Suthipan) | ปกติ
|  | Fluke-179 | 91560315 | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-179 | 91560315 | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-179 | 91560315 | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
|  | Fluke-179 | 91560315 | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2025-09-09 | 2026-09-09 | Head office
151 | FFW Tempcheck | DSTA-40 | 88001127 | Temperature | 0-45 °C | ±0.6 °C | 37-42 °C | N/A | Service admin | 1 ปี | 2025-11-26 | 2026-11-26 | Head office | ปกติ
152 | Hotline Tempcheck | HLTA-40 | 88001703 | Temperature | 0-45 °C | ±0.6 °C | 37-44 °C | N/A | Service admin | 1 ปี | 2025-11-26 | 2026-11-26 | Head office | ปกติ
153 | Incubator/Radiant Warmer Analyzer | INCU II | 64090011 | Temperature | 0-50 °C | ±0.05 °C | 30-42 °C | ±0.3 °C | Service admin | 1 ปี | 2026-02-02 | 2027-02-02 | Head office | ปกติ
154 | Incubator/Radiant Warmer Analyzer | INCU II | 67840014 | Temperature | 0-50 °C | ± 0.5 °C | 30 - 39 °C | ± 1 °C | Service admin | 1 ปี | 2026-03-31 | 2027-03-31 | Head office | ปกติ
155 | Electrical Safety Analyzer | ESA612 | 6845046 | Voltage | 0-300 V | ±2 % of reading | 110-220V | ±2V | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office | ปกติ
|  | ESA612 | 6845046 | Earth Resistance | 0-2Ω | ±2 % of reading | 0.2Ω | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | ESA612 | 6845046 | Equipment Current | 0-20A | ±5 % of reading | 0-18A | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
|  | ESA612 | 6845046 | Leakage current | 0µA-10mA | ±2 % of reading | 0 µA-10mA | ±2 % of reading | Service admin | 1 ปี | 2026-05-29 | 2027-05-29 | Head office
156 | Vital Signs Simulator | Prosim4 | 6662047 | Heart rate | 30-320 BPM | ±1% of Setting | 20-300 BPM | ± 5 BPM | Service admin | 1 ปี | 2026-05-07 | 2027-05-07 | Head office | ปกติ
|  | Prosim4 | 6662047 | Non invasive blood pressure | 10-400mmHg | ±(1% of reading +1mmHg) | 30-290 mmHg | ± 5 mmHg | Service admin | 1 ปี | 2026-05-07 | 2027-05-07 | Head office
|  | Prosim4 | 6662047 | Invasive blood pressure | 0-250 mmHg | ±(1% of reading +1mmHg) | 0-160 mmHg | ± 1 mmHg | Service admin | 1 ปี | 2026-05-07 | 2027-05-07 | Head office
|  | Spot Light | 6662045 | Percent Oxygen | 80-99 % | ±(1+Accuracy of UUT) | 90-99 % | ± 3 % | Service admin | 1 ปี | 2026-04-29 | 2027-04-29 | Head office
157 | Gas Flow Analyzer | VT650 | 6584702 | Flow | 0 to ±200 slpm | ± 2 % | 0 - 180 slpm | ± 10 % of Setting | Service admin | 1 ปี | 2026-05-21 | 2027-05-21 | Head office | ปกติ
|  | VT650 | 6584702 | Pressure | -0.8 to 10 bar | ± 1 % | 0-0.14 bar | ± 0.0026 bar | Service admin | 1 ปี | 2026-05-21 | 2027-05-21 | Head office
|  | VT650 | 6584702 | Oxygen | 0 to 100 % | ± 1 % | 0-100% | ± 3 % of reading | Service admin | 1 ปี | 2026-05-21 | 2027-05-21 | Head office
158 | Gas Flow Analyzer | CITREX H5 | BB106272 | Flow | -300 to 300 l/min | ± 1.9 % | 0 -15 l/min | ± 10 % of Setting | Service admin | 1 ปี | 2026-05-19 | 2027-05-19 | Head office | ปกติ
160 | True RMS Multimeter | Fluke-115 | 4023202WS | AC Curent | 6-10Amp | ± 0.1Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office | ปกติ
|  | Fluke-115 | 4023202WS | AC Voltage | 0-600V | ± 0.5mV | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 4023202WS | DC Current | 0-10Amp | ± 0.013Amp | 0-10Amp | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 4023202WS | DC Voltage | 0-600V | ± 0.06V | 0-600V | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
|  | Fluke-115 | 4023202WS | Resistance | 0-40MΩ | ± 0.04MΩ | 0-40MΩ | N/A | Service admin | 1 ปี | 2026-01-28 | 2027-01-28 | Head office
161 | Steel Weight | Class M1 | M1801082N(ID1)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
162 | Steel Weight | Class M1 | M1801082N(ID3)-5KG | Weight | 0-5kg | ± 80 mg | 0-5000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
163 | Steel Weight | Class M1 | M1801082N(ID4)-1KG | Weight | 0-1kg | ± 16 mg | 0-1000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ
164 | Steel Weight | Class M1 | M1906195N(ID4)-20KG | Weight | 0-20kg | ± 300 mg | 0-20000g | ± 15 % | Service admin | 1 ปี | 2026-02-06 | 2027-02-06 | Head office | ปกติ